High-performance liquid chromatographic quantification of rifampicin in human plasma: method for therapeutic drug monitoring
Trabelsi Sameh 1-2, Eljebari Hanene 1, Charfi Rim 1-2, Salouage Issam1-2, Gaïes Emna1-2 Jebabli Nadia1, Lakhal Mohamed1-2, Klouz Anis 1-2. 
1 Laboratory of Clinical Pharmacology, National Central of Pharmacovigilance, Tunis, Tunisia.
2 Faculty of Medicine, El Manar University, Tunis, Tunisia.
Running Title: Therapeutic drug monitoring of Rifampicin. 
Key words: Rifampicin, liquid chromatography, therapeutic drug monitoring.
Corresponding author: Eljebari Hanene
Adresse : Laboratoire de Pharmacologie Clinique, Centre National de Pharmacovigilance. 
9, Avenue Dr Zouheïr Essafi, 1006 Tunis, Tunisia. 

E- mail: eljebarihanene@yahoo.fr
Tél. : 0021671572300 / 0021671569888

Fax: 00216715713 90 

Abstract: 
A high performance liquid chromatography (HPLC) method has been developed that allows quantification of Rifampicin in human plasma. The method is based on, the precipitation of proteins in human plasma with methanol, and direct injection into HPLC system. A C18 column and a simple mobile phase consisting of 0.05 M dipotassic hydrogen phosphate buffer and acetonitrile (53/47, V/V) and 0.086 % Diethylamin, pH = 4.46. The flow-rate was 0.6 ml/min and the effluent was monitored at 340 nm. Results from the HPLC analyses showed that the assay method is linear in the concentration range 1-40 µg/ml (r²>0.99). The limit of quantification and limit of detection of Rifampicin were respectively 0.632 µg/ml and 0.208 µg/ml. Intraday and interday coefficient of variation and bias were below 10% for all samples, suggesting good precision and accuracy of the method. Recoveries were greater than 90% in a plasma sample volume of 100 µl. The method is being successfully applied to therapeutic drug monitoring of Rifampicin in plasma samples of patients with tuberculosis and staphylococcal infections patients.    

INTRODUCTION 
Rifampicin (RIF), a semi-synthetic macrocyclic complex antibiotic derived from Streptomyces mediterranei, is member of rifamycin class antibiotics (1) used for the treatment of tuberculosis and for other infectious diseases such as staphylococcal infection (2). It is categorised as one of the first line antituberculous agents; however various side effects such as hepatotoxicity, allergic rashes, and lack of appetite, nausea or immunological disturbances have been reported with its administration [2, 3]. The therapeutic drug monitoring (TDM) of Rifampicin is necessary if two months after starting treatment, the cultures are still positive for mycobacterium. TDM is necessary in situation
s where there is a risk of under dosing, especially (i) in the infection with Mycobacterium avium, and the tuberculosis meningitis [4] (ii) in immune deficiency syndrome patients in which the gastrointestinal absorption of antibacterial drugs is reduced and when the incidence of infection by mycobacterium avium increase [5] (iii) Drug interaction (with imidazole antifungal and/or pyrazinamide). The TDM was also necessary in situations where the risk of accumulation and the occurrence of toxic risk are increased particularly in liver dysfunction and renal failure.
TDM of RIF is also useful in case of serious illness associated with gastrointestinal disease, mucovicidose, and diabetes and in child regardless of the clinical situations.
Tuberculosis was considered as a public health problem in Tunisia and the high risks of toxicity with using high Rifampicin doses in antistaphycoccal infection are both justify the necessity of Rifampicin TDM in our laboratory. 
Therefore it was necessary to develop a sensitive, accurate and reproducible analytical method to analyse these samples. Numerous methods have been previously published for the analysis of Rifampicin in biological fluids (plasma, serum, urine,.). The most common described procedures include microbiological assays [6], HPLC thermospray mass spectrometry and electrospray mass spectrometry [7]; HPLC/tandem mass spectrometry method [8 – 12] and HPLC methods [13 – 20].  Most of them needed relatively large volume of samples (>= 0.5 ml) [21- 29] and involve complex or very long extraction procedures, which notably increase the analysis time [21, 23]. The present study was undertaken to develop and validate a simple and rapid HPLC method with UV detection, which could be useful for TDM of Rifampicin in routine setting in patients and to evaluate its pharmacokinetic of Rifampicin.
MATERIALS AND METHODS
Chemicals and reagents 
Lyophilic Rifadine® was provided from Sanofi Aventis (Tunis, Tunisia). The internal standard (IS) Clonazepam (CLONAZEPAM®) was provided by PHARMASCIENCE (Montreal, CANADA). Rifamycin®1UI/100 provided from THEA (France). 
Acetonitrile  HPLC grade, Methanol Lichrosolv, Water Lichrosolv  HPLC grade, K2HPO4 buffer HPLC grade and Diethylamine HPLC grade were obtained from Scharlau (Tunis, Tunisia). Orthophosphoric acid H3PO4 was provided from Prolabo (Tunis, Tunisia). Bidistilled water was prepared in the Laboratory of Clinical Pharmacology of the National Center of Pharmacovigilance (Tunis, Tunisia). 
Preparation solutions 
Rifampicin solution SRF1 (60 mg/ml) was prepared from dilution of 600 mg of Rifampicin powder dissolved in 10 ml of solution provided with powder and stored at -20°C. A  Rifampicin solution SRF2 (600 µg/ml) was prepared by diluting of SRF1 in the bidistilled water. Clonazepam solution (125µg/ml) was prepared in methanol and stored at -20°C in dark. 
Two plasmatic solutions of Rifampicin PRF1 (100 µg/ml) and PRF2 (12 µg/ml) were prepared by diluting in free human plasma. These solutions were freshly prepared before each experiment. They were used for preparation of controls (QCs) and Standards (STD) of Rifampicin (1, 2, 3, 5, 6, 9, 10, 12, 20 and 40 µg/ml) used for validation procedure by diluting in blank human plasma.  
 Blank human plasma for analytical development and validation was obtained from Blood Bank (H.U.C Charles Nicolle, Tunisia). 
HPLC Equipment. 
The HPLC system manufactured by Varian® (Australia) used for the present study consisted of a Varian® Prostar 240 pump, a Prostar 410 Autosampler, Varian® Prostar 325 LC Detector and a Varian®  Prostar 510 Column Oven. For the data acquisition and integration, Galaxie chromatography data system version 1.9.3.2 software operated by Pentium 4HT was used. 

Chromatography conditions 
Rifampicin was separated using a Lichocart cartridge column (150 x 4 mm) filled with Lichrospher 1000 RP-18, having a 5 µm practical size (Merck, France) maintained at 25°C. The samples were eluted in a mobile phase consisting in 0.05 M dipotassic hydrogen phosphate buffer and acetonitrile (53/47, V/V) with 0.086 % Diethylamin, pH = 4.46 at a flow rate of 0.6 ml/min. The mobile phase was filtered through a 0.22 µm filter (Millipore) and degassed under vacuum prior to use. The detector wavelength was set at 340 nm.  
Sample extraction
Aliquots of 100 ml plasma samples (blank, standard or patient sample) and 20 µl of IS solution were vortexed for 1 min. We added 200 µl of methanol to the mixture, re-vortexed for 3 min and let stand 10 min in dark.  Each sample was centrifuged at 4°C (4000 g x 12 min), and then 50 µl of the supernatant was injected into the HPLC column for quantification.
Method Validation   
Our validation study was performed according to FDA guidelines for bioanalytical methods [30].
Linearity

The calibration curves were prepared (n=6)  with five increasing concentrations of Rifampicin  (STD; µg/ml). The final concentrations of Rifampicin in plasma samples were 1, 3, 6, 9, 12, 20 and 40 µg/ml. They were built with the peak areas ratios of each STD and IS, and fitted using linear regression.   
Precision 

Replicate analysis (n=5) of control samples (QCs) at the 4 concentrations (2, 5, 10 and 20 µg/ml) was used for the precision and accuracy determination. All fore concentrations were chosen to encompass the whole range of the calibration curve corresponding to the drug levels frequent in the majority of patients samples. Inter-assay accuracy and precision were determined by repeated analysis performed in three different days. The concentration in each sample was determined using calibration curve prepared on the same day. The precision was calculated as the coefficient of variation (CV %) within a single run (intraday) and between different assays in different days (interdays).  The analytical series were considered valid and accepted only if the percentage of deviation (bias) between theoretical and back-calculated (experimental) concentrations for each calibration level samples were less than 10%.

Accuracy  

Accuracy was measured as the percentage difference between measured value and nominal value, according the equation: Bias (%) = [(Cmeasured – Ctheoretical)/ Ctheoretical] x 100%, where Cmeasured and Ctheoretical are the determined (observed) concentration of Rifampicin with the present method and their theoretical concentrations in the spiked plasma samples, respectively. Extractions were repeated three times for each concentration. 

Determination of limit of detection (LOD) and limit of quantification (LOQ)
LOD was calculated using the equation LOD = 3.3(SD /b) and LOQ was calculated using the equation LOQ = 10(SD /b), where SD is the standard deviation and b is the slope of the calibration curve. 

The aim was to make the bias and CV % less than 10% for all quality controls standards except for the LOQ where less than 20% was accepted. 

Statistical analysis 

Validation data and the equation of the calibration curve were calculated by regression analysis using the software-package Pythagore Biostat (Version 2.0). Data were given as mean  SD. 
Clinical application
The method was successfully applied to the monitoring of Rifampicin in 28 Blood samples provided from patients treated by Rifampicin for tuberculosis or staphylococcal infection disease. Samples were accompanied with patient’s clinical information (Age, body weight, Sex, indication, Rifampicin dose). 
 Samples were collected in lithium heparin tube (3ml) 2h after taking Rifampicin oral dose. Plasma was obtained after blood samples centrifugation at 1400 g x 10 min at 4°C and stored until HPLC analysis at -20°C.        
RESULTS and DISCUSSION 


Method development. 

Different ratios, flow-rates, pHs of mobile phase and column temperatures were studied in order to shorten the retention time of RIF and to improve peak symmetry. However, the optimal assay conditions were found when using ratio of 53/47, V/V (tampon/ACN), and 0.086% of diethylamin, the pH was adjusted to 4.46, the column temperature set at 25°C and the flow-rate at 0.6 ml/min. The peaks were well resolved with λ max at 340 nm. Retention time of RIF in these conditions was 5.2 ± 0.5 min. 
The choice of the internal standard is a crucial step in this work, as we have seen to quantify Rifampicin with an internal calibration technique. The internal standard must be added at the beginning of the extraction and should follow the drug in all stages of extraction. 
It must be well separated from the peak of rifampicin, having a retention time close to the molecule and a total analysis time as short as possible. Several internal standards are described in the literature such as papaverinehydrochloride [16], p-dimethyl [22], Rifapentine [31], Sulindac [32]. We tried two molecules in this study; Clonazepam (5 µg/ml) and Rifamycin (1 UI/100) and the two molecules were detected at retention time of 6.46 and 5.47 min respectively. We have ruled out the Rifamycin because there were impurities accompanying the peak of Rifamycin and we kept Clonazepam as internal standard at the end of this step. The total time for each sample injection was about 10 min. Retention times were approximately 5.2 ± 0.5 min for RIF and 6.3 ± 0.5 for IS for all assays.
Two extraction methods have been described for RIF in literature, a solid-liquid [16], and a Liquid-Liquid [22 -32]. Several extraction solvents have been described; hydrochloric acid 0.1 M [32], acetonitril [16-33]. In our work we have carried out a precipitation of plasmatic proteins using methanol.
Validation 

The specificity of the method was examined by comparing the chromatogram of the blank human plasma with that of the plasma spiked with RIF. Representative chromatogram of blank plasma extracted of RIF is shown in Figure 1A. Representative chromatogram of RIF standard plasma (6 µg) and Clonazepam is shown in Figure 1B. There were no interfering peaks in blank plasma at the retention time of RIF. No carry-over (memory effect) problem was observed in this assay as 50 µl of mobile phase was injected into the HPLC system after analyzing plasma containing RIF. 
Mean regression coefficient (r²) of all RIF calibration curves was 0.9914± 0.008 (n=6) (data not shown). So the method was linear from 
The precision and accuracy of data obtained from controls have been presented in Table 1.  The within day (intra day) CV % of the assay varied from 1.99 % at 10 µg/ml to 8.86 % at 2 µg/ml. The intraday accuracy of estimation in different controls samples was found to be -16.3 % for QC 2 µg/ml, -7.04 % for QC 5 µg/ml, 1.4 % for QC 10 µg/ml and -1.78 % for QC 20 µg/ml. Between day (inter assay) coefficient of variation was 2.78 % for QC 20 µg/ml and 2.9 % or QC2 µg/ml. the inter day accuracy of estimation in the seeded samples varied from -1.66 % to 2.1 % for QC 20 µg/ml and QC 10 µg/ml, respectively. These results suggest good precision and accuracy of the method. 
The recovery of RIF was assessed by comparing the response of five replicates of extracted samples (2, 5, 10 and 20 µg/ml) to the response of pure standard at the same concentration level. The standard specimens were dissolved in bi-distilled water and injected directly into the HPLC system. The recovery of RIF was greater than 90% (data not shown) for all tested concentrations.      
The LOQ of RIF defined as the lowest concentration that could be measured with accuracy and precision, was 0.632 µg/ml. The lowest amount of RIF in the sample, which can be detected but not necessarily quantitated under stated experimental conditions; LOD was 0.208 µg/ml.
Application of the method:

As shown by various chromatograms (chromatogram for the blank sample and chromatogram for 18.7 µg/ml of Rifampicin in plasma from a patient treated for a staphylococcal infection) with 900 mg/day of Rifampicin (figure 2) the present HPLC/UV method seems sufficiently sensitive to be used for maximal concentration (Cmax) Rifampicin plasma level determination
We realized plasma monitoring of Rifampicin in 28 patients (28 samples). All samples were performed 2 hours after taking Rifampicin (Cmax).  
Rifampicin was prescribed for tuberculosis in 17 patients whose average age is 36.7 ±15.5 yrs and average weight is 73.4 ± 18.47 Kg and for staphylococcal infection in 11 patients whose average age is 59.88 ±19.4 yrs and average weight is 67.83 ± 9.21 Kg. Rifampicin doses were varied from 5.36 mg/Kg/day to 12.5 mg/Kg/day, with a mean of 8.83 mg/Kg/day for two indications. Founded concentrations varied from 1.97 µg/ml to 28 µg/ml with a mean of 9.81 µg/ml. In literature target Cmax varied from 4 µg/ml to 24 µg/ml according to the germ, way of action and the dose. A target concentration of 8 µg/ml was desirable. In our population four concentrations were below 4μg/ml, 12 concentrations were ranged between 4 g / ml and 8 µg / ml, and 11 concentrations were up to 8 µg/ml. Study the correlation between the Rifampicin measured concentration and daily dose per kg did not reveal a significant correlation (r² = 0.1798) which justifies the TDM of Rifampicin in our population.       
CONCLUSION
In conclusion, a novel validated method for the quantification of Rifampicin in human plasma was developed using HPLC with UV detection. The assay is accurate, precise and fast. This simple separation method uses isocratic mobile phase conditions making the assay accessible to any laboratory in institute as it can be performed by those unfamiliar to basic HPLC techniques.

REFERENCES

1. Maggi N, Pasqualucci CR, Ballotta R, Sensi P. Rifampicin: a new orally active rifamycin. Chemotherapy. 1966; 11(5):285-292
2. Tsankov N, Angelova I . Rifampin in dermatology.Clin Dermatol 2003 ; 21(1):50-55.

3. Gallieni M, Braidotti P, Cozzolino M, Romagnoli S, Carpani P. Acute tubulo-interstitial nephritis requiring dialysis associated with intermittent rifampicin use: case report. Int J Artif Organs. 1999; 22(7):477-81.
4.  Yew WW. Therapeutic drug monitoring in antituberculosis chemotherapy. Ther Drug Monit1998; 20:469-72.
5. Charles A. Peloquin. Using therapeutic drug monitoring to dose the antimycobacterial grugs. Clinics in Chest Medicine 1997; 18 (1): 79-87
6. Bruzzese T, Rimaroli C, Bonabello A, Mozzi G, Ajay S, Cooverj ND. Pharmacokinetics and tissue distribution of rifametane, a new 3-azinomethyl-rifamycin derivative, in several animal species. Arzneimittelforschung. 2000; 50 (1):60-71.

7. Korfmacher WA, Bloom J, Churchwell MI, Getek TA, Hansen EB Jr, Holder CL, McManus KT.Characterization of three rifamycins via electrospray mass spectrometry and HPLC-thermospray mass spectrometry. J Chromatogr Sci. 1993; 31(12):498-501.

8. Femke de Velde, Jan-Willem C. Alffenaar, A. Mireille A. Wessels, Ben Greijdanus, Donalds R. A. Uges. Simultaneous determination of clarithromycin, rifampicin and their main metabolites in human plasma by liquid chromatography-tandem mass spectrometry. Journal of Chromatography B 2009, 877: 1771-1777. 

9. Stefan O, Jette P, Monica V, Werner S. LC-MS/MS method for the simultaneous determination of clarithromycin, rifampicin and their main metabolites in horse plasma, epithemlial lining fluid and broncho-alveolar cells. Journal of Pharmaceutical and Biomedical analysis 2011; 55: 194-201. 

10. Song SH, Jun SH, Park KU, Yoon Y, Lee JH, Kim JQ, Song J. simultaneous determination of first-line and anti tuberculosis drugs and their major metabolic ratios by liquid chromatography / tandem mass spectrometry. Rapid commun Mass Spectrom 2007; 21 (7): 1331-8. 

11. Ping – Fei Fang, Hua-Lin Cai, Huan-De Li, Rong-Hua Zhu, Qin you Tan, Wei Gao, Ping Xu, Yi-Ping Liu, Wen-Yuan Zhang, Yong-Chang Chen, Feng Zhang. Simultaneous determination of isoniazid, rifampicin, levofloxacin in mouse tissues and plasma by liquid chromatography-tandem mass spectrometry. Journal of chromatography B 2010, 878: 2286-2291. 

12. Lorena B, Antonio D, Francessco G. D. R, Silvia G, Marianna Mchelazzo, Gissi V, Marco S, Marco S, Mauro S, Giovanni D.P. Development and validation of simultaneous extraction procedure for HPLC-MS quantification of daptomycin, amikacin, gentamicin, and rifampicin in human plasma. Anal Bioanal Chem 2010; 396: 791-798. 

13. A Walubo, P Smith, P I Folb Comprehensive assay for pyrazinamide, rifampicin and isoniazid with its hydrazine metabolites in human plasma by column liquid chromatography. Journal of Chromatography B: Biomedical Sciences and Applications 1994; 658 (2):391-396.
14. Moussa LA, Khassouani CE, Soulaymani R, Jana M, Cassanas G, Alric R, Hüe B.Therapeutic isoniazid monitoring using a simple high-performance liquid chromatographic method with ultraviolet detection. J Chromatogr B Analyt Technol Biomed Life Sci 2002; 766 (1):181-187
15. Weber A, Opheim KE, Smith AL, Wong K. High-pressure liquid chromatographic quantification of rifampin and its two major metabolites in urine and serum. Rev Infect Dis. 1983; 5: S433-9. 

16. AL. Allanson, M.M Cotton, J.N.A Tettey, and A.C. Boyter. Determination of rifampicin in human plasma and blood spots by high performance liquid chromatography with UV detection: A potential method for therapeutic drug monitoring. Journal of pharmaceutical and biomedical Analysis 2007; 44: 963-969

17. Ignacio Calleja, Maria J. Blanco-Prieto, Noelia Ruz, Maria Jesus Renedo, Maria Carmen Dios – Vietrez. High-performance liquid-chromatographic determination of Rifampicin in plasma and tissus. Journal Chromatography A. 2004. 1031: 289-294. 

18. Marina S. B, Carlo B, Mateus M.B, Regina H. C. Q, Wilson R. M, Sonia A.C. D, Lidervan P. M, Maria E. C. Q. Rifampicin determination in plasma by stir bar-sorptive extraction and liquid chromatography. Journal of pharmaceutical and Biomedical Analysis 2010; 51: 1078-1083. 

19. Deirder Fox, Robert O’Connor, Patrick Mallon, Gillian Mc Mahon. Simultaneous determination of efaverenz, rifampicin and its metabolite desacetyl rifampicin levels in human plasma. Journal of Pharmaceutical and Biomedical analysis 2011; 56:785-791. 

20. A.K. Hemanth Kumar, I. Chandra, R. Geetha, K. Silambu Chelvi, V. Lalitha, G. Prema. A validated high-performance liquid chromatography method for the determination of rifampicin and desacetyl rifampicin in plasma and urine. Indian Journal of pharmacol 2004; 36 (4): 231-233.
21.  Ratti B, Parenti RR, Toselli A, Zerilli LF. Quantitative assay of rifampicin and its main metabolite 25-desacetylrifampicin in human plasma by reversed-phase high-performance liquid chromatography. J Chromatogr. 1981 9; 225(2):526-31.
22. Chan K. Rifampicin concentrations in cerebrospinal fluid and plasma of the rabbit by high performance liquid chromatography. Methods Find Exp Clin Pharmacol. 1986; 8(12):721-6.

23. Woo J, Wong CL, Teoh R, Chan K. Liquid chromatographic assay for the simultaneous determination of pyrazinamide and rifampicin in serum samples from patients with tuberculous meningitis. J Chromatogr. 1987; 420(1):73-80.

24. Ishii M, Ogata H. Determination of rifampicin and its main metabolites in human plasma by high-performance liquid chromatography. J Chromatogr 1988; 426 (2):412-416.
25. Jamaluddin AB, Sarwar G, Rahim MA, Rahman MK. High-performance liquid chromatographic assay of rifampicin in human serum. J Chromatogr. 1990; 525(2):495-497. 

26. El-Yazigi A, Raines DA. Simultaneous microdetermination of rifampin, deacetylrifampin, isoniazid, and acetylisoniazid in plasma by liquid chromatography with dual electrochemical and spectrophotometric detection. Pharm Res. 1992; 9(6):812-816.
27. Swart KJ, Papgis M. Automated high-performance liquid chromatographic method for the determination of rifampicin in plasma. J Chromatogr. 1992; 593 (1-2):21-24.

28. Walubo A, Smith P, Folb PI. Comprehensive assay for pyrazinamide, rifampicin and isoniazid with its hydrazine metabolites in human plasma by column liquid chromatography. J Chromatogr B Biomed Appl. 1994; 658 (2):391-396.

29. Conte JE Jr, Lin E, Zurlinden E. Liquid chromatographic determination of rifampin in human plasma, bronchoalveolar lavage fluid, and alveolar cells. J Chromatogr Sci. 2000; 38(2):72-76. 

30. http://www.fda.gov/downloads/Drugs/GuidanceComplianceRegulatoryInformation/GuidanGui/UCM070107.pdf. Accessed 5th November 2009
31. Panchagnula R, Sood A, Sharda N, Kaur K, Kaul CL. Determination of Rifampicin and its main metabolite in plasma and urine in presence of pyrazinamide and isoniazid by HPLC method. J Pharm Biomed anal 1999; 18 (6): 1013-1020. 

32. Yau Yi Lau, Gleen D. Hanson, Barbara J. Carl. Determination of Rifampin in human plasma by high-performance liquid chromatography with ultraviolet detection. Journal of Chromatography B 1996; 676: 147-152. 

33. Aparicio I, Bello MA, Callejon M, Guiraum A. Simultanious determination of Rifampicin and sulbactam in mouse plasma by high-performance liquid chromatography. Biomed Chromatogr 2006; 20 (8): 748-752
LEGENDES 
Table1. Precision and accuracy of Rifampicin QCs in human plasma.

n=5 assays in the same day for each concentration; ** n=5 assays per day during 3 days for each concentration. Daily (intra-day) and day to day (inter-day) precisions are represented as mean values ± SD of five different assays for each concentration. 

Figure 1
Representative Chromatograms of blank and Rifampicin spiked plasma (A) Representative chromatogram of blank plasma extracted of Rifampicin (B) Representative chromatogram of Rifampicin Control plasma (Qc 6µg/ml) extracted. 
Figure 2
Representative chromatogram for T2h 10.2 µg/ml of Rifampicin in plasma from a staphylococcal infection patient treated with 900 mg/day of Rifampicin. 
5

